For my for initial prototyping of the camera-based spectrometer path with your U-bend POF LSPR sensor.

The link resolves to a compact handheld diffraction spectroscope (typical gemological style: short aluminum body ~55 mm long, fixed entrance slit, and a transmission diffraction grating). These are functionally very similar to the previous unit you asked about and to the inexpensive pocket spectroscopes commonly used as front-ends in PySpectrometer-style DIY builds.

### How it works in your setup

- Illuminate the entrance slit with light exiting your POF (or a short SMA patch from the sensor end). Your 1 mm core fiber delivers plenty of light, so even a relatively wide fixed slit will produce a usable spectrum.

- Image the spectrum that appears at the eyepiece end with a Raspberry Pi camera (HQ or Module 3 preferred) and a suitable focusing lens.

- Process the image with PySpectrometer (or equivalent) software for intensity extraction, multi-point calibration, averaging, and CSV export. From the calibrated spectrum you can then apply centroid or polynomial fitting for Δλ and intensity metrics.

- Keep the separate DET02AFC + modulated M530F3 path for high-time-resolution intensity kinetics.

### Expected performance for prototyping

- Resolution is educational/gemological grade (typically a few nm). It is sufficient to observe the overall LSPR band shape around 530 nm, confirm that the U-bend is transmitting, and test the full measurement workflow (coupling stability, white-LED illumination, baseline subtraction, peak fitting).

- It is not high enough for precise quantitative tracking of small binding-induced shifts without heavy averaging and fitting, nor does it match a commercial Flame-S with a narrow slit. Treat it as a proof-of-concept tool.

- Mechanical rigidity of both the spectroscope mount and the U-bend holder is still critical to avoid modal noise artifacts.

### Practical notes for your build

- Use a broadband white LED as the source when capturing full spectra (narrowband green clips the shifting LSPR curve).

- Calibrate early with a fluorescent lamp (Hg lines) or known lasers to map pixels to wavelength.

- Start with the bare or baseline sensor to check throughput and stability before functionalized sensors.

- Because these units are inexpensive and widely available, they are a low-risk way to validate the overall approach (fiber → slit → grating → camera → software) before deciding whether to upgrade to a better grating/slit assembly, a TCD1304 linear CCD design, or a commercial mini-spectrometer.

In short, it is a reasonable first prototype tool for home testing of the spectral channel alongside your existing photodetector setup. Once you have it, the highest-value early experiments are fluorescent-lamp calibration and a white-LED spectrum through a plain POF to confirm light levels and line sharpness.